（V1.3）

# Contents

Contents	i

Copyright Declaration	iv

Chapter 1	Safety Tips	1

1.1	General Safety Summary	1

1.2	Safety Terms and Symbols	1

1.3	Terms on Product	2

1.4	Symbols on Product	2

1.5	Product Scrapping	2

Chapter 2	Overview	3

2.1	Brief Introduction to DSO1000B Series	3

2.2	Help System	3

Chapter 3	Getting Started Guide	4

3.1	Installation	4

3.2	Functional Check	4

3.2.1	Power on the oscilloscope	4

3.2.2	Connect the oscilloscope	4

3.2.3	Observe the waveform	5

3.3	Probe Examination	5

3.3.1	Safety	5

3.3.2	Use of Probe Check Wizard	6

3.4	Manual Probe Compensation	6

3.5	Probe Attenuation Setting	7

3.6	Self Calibration	7

Chapter 4	Main Feature Description	8

4.1	Oscilloscope Setup	8

4.2	Trigger	8

4.3	Data Acquisition	9

4.4	Waveform Scaling and Positioning	10

4.5	Waveform Measurement	11

Chapter 5	Basic Operation	12

5.1	DisplayArea	13

5.1.1	XY Format	14

5.2	Horizontal Controls	15

5.2.1	Scan Mode Display (Roll Mode)	17

5.3	Vertical Controls	17

5.3.1	Math FFT	18

5.3.1.1	Setting Time-domain Waveform	19

5.3.1.2	Displaying FFT Spectrum	20

5.3.1.3	Selecting FFT Window	21

5.3.1.4	FFT Aliasing	22

5.3.1.5	Eliminating Aliases	22

5.3.1.6	Magnifying and Positioning FFT Spectrum	23

5.3.1.7	Using Cursors to Measure FFT Spectrum	23

5.4	Trigger Controls	24

5.5	Menu and Option Buttons	29

5.5.1	SAVE/RECALL	29

5.5.2	MEASURE	30

5.5.3	ACQUIRE	31

5.5.4	UTILITY	33

5.5.5	CURSOR	33

5.5.6	DISPLAY	34

5.6	Fast Action Buttons	35

5.6.1	AUTO	35

5.7	Signal Connectors	36

Chapter 6	Application Examples	38

6.1	Example 1: Taking Simple Measurements	38

6.2	Example 2: Taking Cursor Measurements	40

6.3	Example 3: Analyzing Input Signals to Eliminate Random Noise	42

6.4	Example 4: Capturing Single-shot Signal	44

6.5	Example 5: Using X-Y Mode	45

6.6	Example 6: Triggering on Pulse Width	46

6.7	Example 7: Triggering on Video Signal	47

6.8	Example 8: Using Slope Trigger to Capture Particular Slope Signal	49

6.9	Example 9: Using Overtime Trigger to Measure Long Pulse Signal	49

6.10	Example 10: Using Math Functions to Analyze Waveforms	50

6.11	Example 11: Measuring Data Propagation Delay	51

6.12	Example 12: Setting Net IP	53

Chapter 7	Multimeter	54

Chapter 8	Troubleshooting	64

8.1	Problem Settlement	64

Chapter 9	Specifications	65

9.1	Technical Specifications	65

9.2	Accessories	71

Chapter 10	Services and Support	72

Chapter 11	General Care and Cleaning	73

11.1	General Care	73

11.2	Cleaning	73

Appendix A  Harmful and Poisonous Substances or Elements	74

# Copyright Declaration

All rights reserved; no part of this document may be reproduced or transmitted in any form or by any means, electronic or mechanical, without prior written permission from Qingdao Hantek Electronic Co., Ltd (hereinafter referred to as ‘Hantek’).

Hantek reserves all rights to modify this document without prior notice. Please contact Hantek for the latest version of this document before placing an order.

Hantek has made every effort to ensure the accuracy of this document but does not guarantee the absence of errors. Moreover, Hantek assumes no responsibility in obtaining permission and authorization of any third party patent, copyright or product involved in relation to the use of this document.

# Safety Tips

## General Safety Summary

Read the following safety precautions to avoid injury and prevent damage to this product or any products connected to it. To evade potential hazards, use this product only as specified.

Only qualified personnel should perform maintenance.

Avoid fire or personal injury.

Connect and disconnect properly. Connect a probe with the oscilloscope before it is connected to measured circuits; disconnect the probe from the oscilloscope after it is disconnected from measured circuits.

Connect the probe in a right way. The probe ground lead is at ground potential. Do not connect the ground lead to an elevated voltage.

Check all terminal ratings. To avoid fire or shock hazard, check all ratings and markings on the product. Refer to the product manual for detailed information about ratings before making connections to the product.

Do not operate without covers. Do not operate this product with covers or panels removed.

Avoid exposed circuitry. Do not touch exposed connections and components when power is present.

Do not operate with suspected failures. If you suspect there is damage to this product, have it inspected by qualified service personnel.

Assure good ventilation.

Do not operate in wet/damp environments.

Do not operate in an explosive atmosphere.

Keep product surfaces clean and dry.

## Safety Terms and Symbols

The following terms may appear in this manual:

WARNING. Warning statements point out conditions or practices that could result in injury or loss of life.

CAUTION. Caution statements identify conditions or practices that could result in damage to this product or other property.

## Terms on Product

The following terms may appear on the product:

DANGER indicates an injury hazard immediately accessible as you read the marking.

WARNING indicates an injury hazard not immediately accessible as you read the marking.

CAUTION indicates a possible hazard to this product or other property.

## Symbols on Product

The following symbols may appear on the product:

## Product Scrapping

Device Recycling

We need extract and utilize natural resources to produce this device. If you do not reclaim the device in a proper way, some substances it contains may become harmful or poisonous to environments or human bodies. To avoid them being released outside and to minimize the waste of natural resources, we suggest you reasonably call back this device to ensure proper recovery and recycling of most materials within it.

# Overview

## Brief Introduction to DSO1000B Series

| Model | Channels | Bandwidth | Sample Rate | LCD |
| --- | --- | --- | --- | --- |
| DSO1062B | 2 | 60MHz | 1GS/s | 7 inch color |
| DSO1102B | 2 | 100MHz | 1GS/s | 7 inch color |
| DSO1202B | 2 | 200MHz | 1GS/s | 7 inch color |

Table 2-1 Model List of DSO1000B Series

DSO1000B Series oscilloscopes cover the bandwidths from 60MHz to 200MHz, and provide the real-time and equivalent sample rates respectively up to 1GSa/s and 25GSa/s. In addition, they have maximum 1M memory depth for better observation of the waveform details, and 5.7 inch color TFT LCD as well as WINDOWS-style interfaces and menus for easy operation.

What’s more, the plenty menu information and the easy-to-operate buttons allow you to gain information as much as possible in measurement; the multifunctional and powerful shortcut keys help you save a lot of time in operation; the Autoset function lets you detect sine and square waves automatically; the Probe Check Wizard guides you to adjust the probe compensation and set the Probe option attenuation factor. By using these the methods the oscilloscope provides (context-sensitive, hyperlinks, and an index), you may master all operations on the device in quite a short time so as to greatly improve your efficiency in production and development.

## Help System

This oscilloscope has a Help system with topics covering all of its features. You can use the Help system to display several kinds of information:

- General information about understanding and using the oscilloscope, such as Using the Menu System.

- Information about specific menus and controls, such as the Vertical Position Control.

- Advice to problems you may come across while using an oscilloscope, such as Reducing Noise.

# Getting Started Guide

## Installation

To keep proper ventilation of the oscilloscope in operation, leave a space of more than 5cm away from the top and the two sides of the product.

## Functional Check

Follow the steps below to perform a quick functional check to your oscilloscope.

### Power on the oscilloscope

Press the ON/OFF button. The default Probe option attenuation setting is 1X.

The default probe parameter

### Connect the oscilloscope

Set the switch on the probe to 1X and connect the probe to Channel 1 on the oscilloscope. First, align the slot in the probe connector with the protuberance on the CH1 BNC and push to connect; then, turn to right to lock the probe in place; after that, connect the probe tip and reference lead to the PROBE COMP connectors. There is a mark on the panel: Probe COMP ~2V@1KHz.

### Observe the waveform

Press the AUTO button and you should see within a few seconds a square wave of about 2V peak-to-peak at 1kHz in the display. Press the CH1 MENU button and remove Channel 1. Push the CH2 MENU button and repeat Step 2 and Step 3 to observe Channel 2.

## Probe Examination

### Safety

When using the probe, keep your fingers behind the guard on the probe body to avoid electric shock. Do not touch metallic portions of the probe head while it is connected to a voltage source. Connect the probe to the oscilloscope and connect the ground terminal to ground before you start any measurements.

### Use of Probe Check Wizard

Every time you connect a probe to an input channel, you should use the probe check wizard to verify that this probe is operating correctly. There are two ways to do this:

- Use the vertical menu (for example, push the CH1 MENU button) to set the Probe option attenuation factor.

- Press UTILITY ->F5->F5->F2 to use the Probe Check Wizard and configure the probe option attenuation factor properly following menu prompts.

## Manual Probe Compensation

Upon the first connection of a probe and an input channel, you should manually perform this adjustment to match the probe to the input channel. Uncompensated or miscompensated probes may lead to errors or faults in measurement. To adjust the probe compensation, follow the steps below.

- Set the Probe option attenuation in the channel menu to 10X. Set the switch on the probe to 10X and connect the probe to Channel 1 on the oscilloscope. If you use the probe hook-tip, ensure it is firmly inserted onto the probe. Attach the probe tip to the PROBE COMP ~5V@1KHz connector and the reference lead to the PROBE COMP Ground connector. Display the channel and then press the AUTO button.

- Check the shape of the displayed waveform.

- If necessary, use a nonmetallic screwdriver to adjust the variable capacity of your probe until the shape of the waveform turns to be the same as the above figure. Repeat this step as necessary. See the figure below for the way of adjustment.

## Probe Attenuation Setting

Probes are of various attenuation factors which affect the vertical scale of the signal. The Probe Check function is used to verify if the Probe attenuation option matches the attenuation of the probe.

As an alternative method to Probe Check, you can push a vertical menu button (such as the CH 1 MENU button) and select the Probe option that matches the attenuation factor of your probe.

Make sure that the Attenuation switch on the probe matches the Probe option in the oscilloscope. Switch settings are 1X and 10X.

When the Attenuation switch is set to 1X, the probe limits the bandwidth of the oscilloscope to 6MHz. To use the full bandwidth of the oscilloscope, be sure to set the switch to 10X.

## Self Calibration

The self calibration routine helps optimize the oscilloscope signal path for maximum measurement accuracy. You can run the routine at any time but should always run it if the ambient temperature changes by 5℃ or more. For a more accurate calibration, please power on the oscilloscope and wait for 20 minutes until it has adequately warmed up.

To compensate the signal path, disconnect any probes or cables from the front-panel input connectors. Then, push the UTILITY button, select the Do Self Cal option and follow the directions on the screen.

.

# Main Feature Description

This chapter provides some general information that you need to learn before using an oscilloscope. It contains:

1. Oscilloscope setup

2. Trigger

3. Data acquisition

4. Waveform scaling and positioning

5. Waveform measurement

## Oscilloscope Setup

While operating the oscilloscope, you may often use thr feature: Autoset.

Autoset: This function can be used to adjust the horizontal and vertical scales of the oscilloscope automatically and set the trigger coupling, type, position, slope, level and mode, etc., to acquire a stable waveform display.

## Trigger

The trigger determines when the oscilloscope begins to acquire data and display a waveform. Once a trigger is properly set up, the oscilloscope can convert unstable displays or blank screens to meaningful waveforms. Here introduce some basic concepts about trigger.

Trigger Source: The trigger can be generated with multiple sources. The most common one is the input channel (alternative between CH1 and CH2). Whether the input signal is displayed or not, it can trigger normal operations.

Trigger Type: The oscilloscope has six types of triggers: Edge, Video, Pulse Width, Slope, Overtime and Swap.

- Edge Trigger uses the analog or digital test circuits for triggering. It happens when the input trigger source crosses a specified level in a specified direction.

- Video Trigger performs a field or line trigger through standard video signals.

- Pulse Width Trigger can trigger normal or abnormal pulses that meet trigger conditions.

- Slope Trigger uses the rise and fall times on the edge of signal for triggering.

- Overtime Trigger happens after the edge of signal reaches the set time.

- Swap Trigger, as a feature of analog oscilloscopes, gives stable displays of signals at two different frequencies. Mainly it uses a specific frequency to switch between two analog channels CH1 and CH2 so that the channels will generate swap trigger signals through the trigger circuitry.

Trigger Mode: You can select the Auto or Normal mode to define how the oscilloscope acquires data when it does not detect a trigger condition. Auto Mode performs the acquisition freely in absence of valid trigger. It allows the generation of untriggered waveforms with the time base set to 80ms/div or slower. Normal Mode updates the displayed waveforms only when the oscilloscope detects a valid trigger condition. Before this update, the oscilloscope still displays the old waveforms. This mode shall be used when you want to only view the effectively triggered waveforms. In this mode, the oscilloscope displays waveforms only after the first trigger. To perform a single sequence acquisition, push the SINGLE SEQ button.

Trigger Coupling: Trigger Coupling determines which part of the signal will be delivered to the trigger circuit. This can help to obtain a stable display of the waveform. To use trigger coupling, push the TRIG MENU button, select an Edge or Pulse trigger, and then select a Coupling option.

Trigger Position: The horizontal position control establishes the time between the trigger position and the screen center.

Slope and Level: The Slope and Level controls help to define the trigger. The Slope option determines whether the trigger point is on the rising or falling edge of a signal. To perform the trigger slope control, press the TRIG MENU button, select an Edge trigger, and use the Slope button to select rising or falling. The LEVEL button controls the trigger point is on which position of the edge.

## Data Acquisition

When you acquire an analog signal, the oscilloscope will convert it into a digital one.  There are two kinds of acquisition: Real-time acquisition and Equivalent acquisition. The real-time acquisition has three modes: Normal, Peak Detect, and Average. The acquisition rate is affected by the setting of time base.

Normal: In this acquisition mode, the oscilloscope samples the signal in evenly spaced intervals to establish the waveform. This mode accurately represents signals in most time. However, it does not acquire rapid variations in the analog signal that may occur between two samples, which can result in aliasing and may cause narrow pulses to be missed. In such cases, you should use the Peak Detect mode to acquire data.

Peak Detect: In this acquisition mode, the oscilloscope gets the maximum and minimum values of the input signal over each sample interval and uses these values to display the waveform. In this way, the oscilloscope can acquire and display those narrow pulses that may have otherwise been missed in Normal mode. However, noise will appear to be higher in this mode.

Average: In this acquisition mode, the oscilloscope acquires several waveforms, averages them, and displays the resulting waveform. You can use this mode to reduce random noise.

Equivalent Acquisition: This kind of acquisition can be utilized for periodic signals. In case the acquisition rate is too low when using the real-time acquisition, the oscilloscope will use a fixed rate to acquire data with a stationary tiny delay after each acquisition of a frame of data. After repeating this acquisition for N times, the oscilloscope will arrange the acquired N frames of data by time to make up a new frame of data. Then the waveform can be recovered. The number of times N is related to the equivalent acquisition rate.

Time Base: The oscilloscope digitizes waveforms by acquiring the value of an input signal at discrete points. The time base helps to control how often the values are digitized. Use the TIME/DIV button to adjust the time base to a horizontal scale that suits your purpose.

## Waveform Scaling and Positioning

The display of waveforms on the screen can be changed by adjusting their scale and position. Once the scale changes, the waveform display will increase or decrease in size. Once the position changes, the waveform will move up, down, right, or left.

The channel reference indicator (located on the left of the graticule) identifies each waveform on the screen. It points to the ground level of the waveform record.

Vertical Scale and Position: The vertical position of a waveform can be changed by moving it up and down on the screen. To compare data, you may align a waveform over another. When you push the VOLTS button to change the vertical scale of a waveform, the waveform display will contract or expand vertically to the ground level.

Horizontal Scale and Position: Pretrigger Information

You can adjust the HORIZONTAL POSITION control to view waveform data before the trigger, after the trigger, or some of each. When you change the horizontal position of a waveform, you are actually changing the time between the trigger position and the screen center.

For example, if you want to find out the cause of a glitch in your test circuit, you should trigger on the glitch and make the pretrigger period long enough to capture data before the glitch. Then you can analyze the pretrigger data and perhaps find the cause. You are allowed to change the horizontal scale of all the waveforms by clicking the TIME/DIV button. For example, you may want to see just one cycle of a waveform to measure the overshoot on its rising edge. The oscilloscope shows the horizontal scale as time per division in the scale readout. Since all active waveforms use the same time base, the oscilloscope only displays one value for all the active channels.

## Waveform Measurement

The oscilloscope displays graphs of voltage versus time and can help to measure the displayed waveform. There are several ways to take measurements, using the graticule, the cursors or performing an automatic measurement.

Graticule: This method allows you to make a quick, visual estimate and take a simple measurement through the graticule divisions and the scale factor.

For example, you can take simple measurements by counting the major and minor graticule divisions involved and multiplying by the scale factor. If you counted 6 major vertical graticule divisions between the minimum and maximum values of a waveform and knew you had a scale factor of 50mV/division, you could easily calculate your peak-to-peak voltage as follows:

6 divisions x 50mV/division = 300mV.

Cursor: This method allows you to take measurements by moving the cursors. Cursors always appear in pairs and the displayed readouts are just their measured values. There are two kinds of cursors: Amplitude Cursor and Time Cursor. The amplitude cursor appear as a horizontal broken line, measuring the vertical parameters. The time cursor appear as a vertical broken line, measuring the horizontal parameters.

When using cursors, please make sure to set the Source to the waveform that you want to measure on the screen. To use cursors, push the CURSOR button.

Automatic Measurement: The oscilloscope performs all the calculations automatically in this mode. As this measurement uses the waveform record points, it is more precise than the graticule and cursor measurements. Automatic measurements show the measurement results by readouts which are periodically updated with the new data acquired by the oscilloscope.

# Basic Operation

The front panel of the oscilloscope is divided into several functional areas. In this chapter we will give a quick overview of all control buttons on the front panel as well as the displayed information on the screen and relative testing operations. The figure below illustrates the front panel of the DSO1000B series digital oscilloscope.

Front Panel of DSO1000B Series

## DisplayArea

- Display Format:

- Acquisition Mode: Normal, Peak Detect or Average

- Trigger Status:

The oscilloscope is acquiring pretriggered data.

All pretriggered data have been acquired and the oscilloscope is ready to accept a trigger.

T    The oscilloscope has detected a trigger and is acquiring the posttrigger information.

The oscilloscope works in auto mode and is acquiring waveforms in the absence of triggers.

The oscilloscope is acquiring and displaying waveform data continuously in scan mode.

●    The oscilloscope has stopped acquiring waveform data.

S    The oscilloscope has finished a single sequence acquisition.

- Tool Icon:

: If this icon lights up, it means the keyboard of the oscilloscope is locked by the host computer via USB control.

: If this icon lights up, it means the USB disk has been connected.

: This icon lights up only when the USB slave interface is connected with the computer.

- Readout shows main time base setting.

- Main Time Base Window

- Display of window’s position in data memory and data length.

- Window Time Base

- Operating Menu shows different information for different function keys.

- Icon indicates channel coupling.

- Level Range.

- Icon indicates whether the waveform is inverted or not.

- 20M Bandwidth Limit. If this icon lights up, it means the bandwidth limit is enabled, otherwise disabled.

- Trigger Type:

: Edge trigger on the rising edge.

: Edge trigger on the falling edge.

: Video trigger with line synchronization.

: Video trigger with field synchronization.

: Pulse Width trigger, positive polarity.

: Pulse Width trigger, negative polarity.

- Trig Level.

- Channel Marker

- Window displays waveform.

### XY Format

The XY format is used to analyze phase differences, such as those represented by Lissajous patterns. The format plots the voltage on CH1 against the voltage on CH2, where CH1 is the horizontal axis and CH2 is the vertical axis. The oscilloscope uses the untriggered Normal acquisition mode and displays data as dots. The sampling rate is fixed at 1 MS/s.

The oscilloscope can acquire waveforms in YT format at any sampling rate. You may view the same waveform in XY format. To perform this operation, stop the acquisition and change the display format to XY.

The table below shows how to operate some controls in XY format.

| Controls | Usable or not in XY format |
| --- | --- |
| CH1 VOLTS and VERTICAL POSITION controls | Set the horizontal scale and position |
| CH2 VOLTS and VERTICAL POSITION controls | Continuously set the vertical scale and position |
| Reference or Math | Unusable |
| Cursors | Unusable |
| Autoset (display format reset to YT) | Unusable |
| Time base controls | Unusable |
| Trigger controls | Unusable |

## Horizontal Controls

Use the horizontal controls to change the horizontal scale and position of waveforms. The horizontal position readout shows the time represented by the center of the screen, using the trigger time as zero. When you change the horizontal scale, the waveform will expand or contract to the screen center. The readout near the upper right of the screen shows the current horizontal position in second. M represents ‘Main Time Base’, and W indicates ‘Window Time Base’. The oscilloscope also has an arrow icon at the top of the graticule to indicate the horizontal position.

1. HORIZONTAL POSITION BAR: Used to control the trigger position against the screen center.

2. Each option in HORI MENU is described as follows.

| Options | Settings | Comments |
| --- | --- | --- |
| Window Control | Major Window Minor Window | Selects the major or minor window in dual-window mode. The window is highlighted once selected. Press this option button in single-window mode to enter the daul-window mode. |
| Mark | Right arrow Left arrow Set/Clear Clear All | This function is usable only in dual-window mode. It sets marks at some waveform record locations that users are interested in, and searches for these marks by right and left arrows. Then it positions the window to this mark for further observation. |
| Holdoff | None | Select this menu and click the up and down Arrow keys to adjust the trigger holdoff time within the range of 100ns-10s. |
| Autoplay | None | This function is usable in dual-window mode. Push this menu button and auto move it from left to right at a specified speed. In the expanded window will display corresponding waveforms until it stops once reaching the rightmost side of the major scan window |

Single-window Mode

Dual-window Mode (Full Screen)

3.TIME/DIV: Used to change the horizontal time scale so as to magnify or compress the waveform horizontally. If the waveform acquisition is stopped (by using the RUN/STOP button), the TIME /DIV control will expand or compress the waveform. In dual-window mode, push button F1 to select major or minor window. When the major window is selected, the F1 button provides the same functions as it provides in single-mode window. When the minor window is selected, press TIME/DIV button to scale the waveform whose magnification is up to 1000.

### Scan Mode Display (Roll Mode)

With the TIME/DIV control set to 80ms/div or slower and the trigger mode set to Auto, the oscilloscope works in the scan acquisition mode. In this mode, the waveform display is updated from left to right without any trigger or horizontal position control.

## Vertical Controls

Vertical controls can be used to display and remove waveforms, adjust vertical scale and position, set input parameters and perform math calculations. Each channel has a separate vertical menu to set. See below for menu description.

- VERTICAL POSITION Bar: Move the channel waveform up and down on the screen. In dual-window mode, move the waveforms in both windows at the same time in a same direction. Two channels correspond to two bars.

- Menu (CH1, CH2): Display vertical menu options; turn on or off the display of channel waveforms.

| Options | Settings | Comments |
| --- | --- | --- |
| Coupling | DC AC Ground | DC passes both DC and AC components of the input signal. AC blocks the DC component of the input signal and attenuates signals below 10Hz. Ground disconnects the input signal. |
| 20MHz Bandwidth Limit | Unlimited Limited | Limits the bandwidth to reduce display noise; filters the signal to eliminate noise and other unnecessary HF components. |
| VOLTS | Coarse Fine | Selects the resolution of the VOLTS bar. Coarse defines a 1-2-5 sequence. Fine changes the resolution to small steps between the Coarse settings. |
| Probe Attenuation | 1X 10X 100X 1000X | Selects a value according to the probe attenuation factor so as to ensure correct vertical readouts. Reduce bandwidth to 6MHz when using a 1X probe. |
| Invert | Off On | Inverts the waveform relative to the reference level. |

Ground Coupling

Used to display a zero-volt waveform. Internally, the channel input is connected with a zero-volt reference level.

Fine Resolution

In the fine resolution setting, the vertical scale readout displays the actual VOLTS setting. The vertical scale changes only after you adjust the VOLTS control and set to coarse.

Remove Waveform Display

To remove a waveform from the screen, first push the menu button to display the vertical menu, then push again to remove the waveform. A channel waveform which is unnecessary to be displayed can be used as a trigger source or for math operations.

- VOLTS

Control the oscilloscope to magnify or attenuate the source signal of the channel waveform. The vertical size of the display on the screen will change (increase or decrease) to the ground level. Also you may use the key F3 to switch between coarse and fine.

- MATH MENU: Display the waveform math operations. See the table below for details.

The MATH menu contains source options for all math operations.

| Operations | Source Options | Comments |
| --- | --- | --- |
| ＋ | CH1+CH2 | Add Channel 1 to Channel 2. |
| － | CH1-CH2 | Subtract the Channel 2 waveform from the Channel 1 waveform. |
|  | CH2-CH1 | Subtract the Channel 1 waveform from the Channel 2 waveform. |
| FFT | CH1 or CH2 | Three types of window available for selection: Hanning, Flattop, Rectangular. |
|  |  | Zoom: Use the FFT Zoom button to adjust the window size. Scale: x1, x2, x5, x10. |

Note: All selected menus are highlighted in orange.

### Math FFT

This chapter elaborates how to use the Math FFT (Fast Fourier Transform). You can use the Math FFT mode to convert a time-domain (YT) signal into its frequency components (spectrum), and to observe the following types of signals:

- Analyze harmonics in power cords;

- Measure harmonic content and distortion in systems;

- Characterize noise in DC power supplies;

- Test impulse response of filters and systems;

- Analyze vibration.

To use the Math FFT mode, perform the following tasks:

- Set the source (time-domain) waveform;

- Display the FFT spectrum;

- Choose a type of FFT window;

- Adjust the sample rate to display the fundamental frequency and harmonics without aliasing;

- Use zoom controls to magnify the spectrum;

- Use cursors to measure the spectrum.

#### Setting Time-domain Waveform

It is necessary to set the time-domain (YT) waveform before using the FFT mode. Follow the steps below.

- Push the AUTO button to display a YT waveform.

- Click the Vertical Position key to vertically move the YT waveform to the center (zero division) so as to ensure the FFT will display a true DC value.

- Click the Horizontal Position key to position the part of the YT waveform to be analyzed in the center eight divisions of the screen. The oscilloscope uses the 2048 center points of the time-domain waveform to calculate the FFT spectrum.

- Click the VOLTS Key to ensure the entire waveform remains on the screen. If the entire waveform is invisible, the oscilloscope may display wrong FFT results by adding high-frequency components.

- Click the TIME/DIV key to provide the resolution you need in the FFT spectrum.

- If possible, set the oscilloscope to display multiple signal cycles.

If you Click the TIME/DIV key to select a faster setting (fewer cycles), the FFT spectrum will display a larger frequency range and reduce the possibility of FFT aliasing.

To set the FFT display, follow the steps below.

- Push the M/R button;

- Set the Operation option to FFT;

- Select the Math FFT Source channel.

In many situations, the oscilloscope can also generate a useful FFT spectrum despite the YT waveform not being triggered. This is especially true if the signal is periodic or random (such as noise).

Note: You should trigger and position transient or burst waveforms as close as possible to the screen center.

Nyquist Frequency

The highest frequency that any real-time digital oscilloscope can measure without errors is half of the sample rate, which is called the Nyquist frequency. Frequency information beyond the Nyquist frequency is undersampled which brings about the FFT aliasing. The math function can convert the center 2048 points of the time-domain waveform to an FFT spectrum. The resulting FFT spectrum contains 1024 points from DC (0Hz) to the Nyquist frequency. Usually, the screen compresses the FFT spectrum horizontally to 250 points, but you can use the FFT Zoom function to expand the FFT spectrum so that you can clearly view the frequency components at each of the 1024 data points in the FFT spectrum.

Note: The oscilloscope’s vertical response is a little bit larger than its bandwidth (60MHz, 100MHz or 200MHz, depending on the model; or 20MHz when the Bandwidth Limit option is set to Limited). Therefore, the FFT spectrum can display valid frequency information above the oscilloscope bandwidth. However, the amplitude information near or above the bandwidth will not be accurate.

#### Displaying FFT Spectrum

Push the MATH MENU button to display the Math menu. Use the options to select the Source channel, the Window algorithm and the FFT Zoom factor. Only one FFT spectrum can be displayed at a time.

| Math FFT Options | Settings | Comments |
| --- | --- | --- |
| Source | CH1, CH2 | Choose a channel to be the FFT source. |
| Window | Hanning, Flattop, Rectangular | Select a type of the FFT window. For more information, refer to Section 5.3.1.3. |
| FFT Zoom | X1, X2, X5, X10 | Change the horizontal magnification of the FFT display. For detailed information, refer to Section 5.3.1.6. |

- Frequency at the center graticule line

- Vertical scale in dB per division (0dB=1VRMS)

- Horizontal scale in frequency per division

- Sample rate in number of samples per second

- FFT window type

#### Selecting FFT Window

Using windows can eliminate the spectral leakage in the FFT spectrum. The FFT algorithm assumes that the YT waveform repeats all the time. When the number of cycles is integral (1, 2, 3 ...), the YT waveform starts and ends at the same amplitude and there are no discontinuities in the signal shape.

If the number of cycles is nonintegral, the YT waveform starts and ends at different amplitudes and transitions between the start and end points will cause discontinuities in the signal that introduces high-frequency transients.

Applying a window to the YT waveform changes the waveform so that the start and stop values are close to each other, which reduces the discontinuities.

The Math FFT function has three FFT Window options. There is a trade-off between frequency resolution and amplitude accuracy for each type of window. You shall determine which one to choose according to the object you want to measure and the source signal characteristics.

| Window | Measurement | Characteristics |
| --- | --- | --- |
| Hanning | Periodic Waveform | Better frequency, poorer amplitude accuracy than Flattop |
| Flattop | Periodic Waveform | Better amplitude, poorer frequency accuracy than Hanning |
| Rectangular | Pulse or Transient Waveform | Special-purpose window applicable to discontinuous waveforms. This is actually the same as no windows. |

#### FFT Aliasing

Problems occur when the time-domain waveform acquired by the oscilloscope contains frequency components higher than the Nyquist frequency. The frequency components above the Nyquist frequency will be undersampled and displayed as lower frequency components that ‘fold back’ from the Nyquist frequency. These erroneous components are called aliases.

#### Eliminating Aliases

To eliminate aliases, use the following methods.

- Click the TIME/DIV key to set a faster sample rate. Because the Nyquist frequency increases as you increase the sample rate, the aliased frequency components will be displayed correct. If too many frequency components appear on the screen, you may use the FFT Zoom option to magnify the FFT spectrum.

- If there is no need to observe the frequency components above 20MHz, set the Bandwidth Limit option to Limited.

- Filter the signal input from outside and limit the bandwidth of the source waveform to lower than the Nyquist frequency.

- Identify and ignore the aliased frequencies.

- Use zoom controls and cursors to magnify and measure the FFT spectrum.

#### Magnifying and Positioning FFT Spectrum

You may scale the FFT spectrum and use cursors to measure it through the FFT Zoom option which enables the horizontal magnification. To vertically magnify the spectrum, use the vertical controls.

Horizontal Zoom and Position

You can use the FFT Zoom option to magnify the FFT spectrum horizontally without changing the sample rate. The available zoom factors are X1(default), X2, X5 and X10. When the zoom factor is set to X1 and the waveform is located at the center graticule, the left graticule line is at 0Hz and the right is at the Nyquist frequency.

You magnifies the FFT spectrum to the center graticule line when you change the zoom factor. That is, the axis for horizontal magnification is the center graticule line. Click the Horizontal Position Key to move the FFT spectrum to the right.

Vertical Zoom and Position

When the FFT spectrum is being displayed, the channel vertical keys become the zoom and position controls corresponding to their respective channels. The VOLTS key provides the following zoom factors: X1(default), X2, X5 and X10. The FFT spectrum is magnified vertically to the marker M (math waveform reference point on the left edge of the screen). Click the Vertical Position key to move up the spectrum.

#### Using Cursors to Measure FFT Spectrum

You may use cursors to take two measurements on the FFT spectrum: amplitude (in dB) and frequency (in Hz). Amplitude is referenced to 0db that equals 1VRMS here. You may use cursors to measure at any zoom factor.

Push the CURSOR button, choose the Source option and then select Math. Press the Type option button to select between Amplitude and Frequency. Click the SELECT CURSOR option to choose a cursor. Then move Cursor S and Cursor E. Use the horizontal cursor to measure the amplitude and the vertical cursor to measure the frequency. Now the display at the DELTA menu is just the measured value, and the values at Cursor S and Cursor E.

Delta is the absolute value of Cursor S minus Cursor E.

## Trigger Controls

The trigger can be defined through the Trigger Menu. There are six types of trigger: Edge, Video, Pulse Width, Swap, Slope and Overtime. Refer to the following tables to find a different set of options for each type of trigger.

TRIG MENU

Push this button to display trigger menus. The edge trigger is in common use. See the table below for details.

| Options | Settings | Comments |
| --- | --- | --- |
| Trigger Type |  |  |
| Edge Video Pulse Slope Swap Overtime |  | By default the oscilloscope uses the edge trigger which triggers the oscilloscope on the rising or falling edge of the input signal when it crosses the trigger level (threshold). |
| Source | CH1 CH2 AC Line | Select the input source as the trigger signal. CH1, CH2: No matter the waveform is displayed or not, a certain channel will be triggered. AC Line: Uses a signal derived from the power cord as the trigger source. |
| Mode | Auto Normal | Select a trigger mode. By default, the oscilloscope uses the Auto mode. In this mode, the oscilloscope is forced to trigger when it does not detect a trigger within a certain amount of time based on the TIME/DIV setting. The oscilloscope goes into the scan mode at 80ms/div or slower time base settings. In the Normal mode, the oscilloscope updates the display only when it detects a valid trigger condition. New waveforms are not displayed until they replace old ones. Use this mode to just view valid triggered waveforms. Only after the first trigger does the display appear. |
| Coupling | AC DC HF Reject LF Reject | Select the components of the trigger signal applied to the trigger circuitry. AC: Blocks DC components and attenuates signals below 10Hz. DC: Passes all components of the signal. HF Reject: Attenuates the high-frequency components above 80kHz. LF Reject: Blocks DC components and attenuates the low-frequency components below 8kHz. |

NOTE: Trigger coupling only affects the signal passed through the trigger system. It does not affect the bandwidth or coupling of the signal displayed on the screen.

Video Trigger

| Options | Settings | Comments |
| --- | --- | --- |
| Video |  | With Video highlighted, an NTSC, PAL or SECAM standard video signal will be triggered. The trigger coupling is preset to AC. |
| Source | CH1 CH2 | Select the input source as the trigger signal. |
| Polarity | Normal Inverted | Normal: Triggers on the negative edge of the sync pulse. Inverted: Triggers on the positive edge of the sync pulse. |
| Sync | All Lines Line Number Odd Field Even Field All Fields | Choose a proper video sync. When selecting Line Number for the Sync option, you may use the User Select it to specify a line number. |
| Standard | NTSC PAL/SECAM | Choose a video standard for sync and line number count. |

Note: When you choose Normal Polarity, the trigger always occurs on negative-going sync pulses. If the video signal contains positive-going sync pulses, use the Inverted Polarity option.

Pulse Width Trigger

You can use it to trigger on aberrant pulses.

| Options | Settings | Comments |
| --- | --- | --- |
| Pulse |  | With Pulse highlighted, the trigger occurs on pulses that meet the trigger condition (defined by the Source, When and Set Pulse Width options). |
| Source | CH1 CH2 | Select the input source as the trigger signal. |
| When | = ≠ < > | Select the trigger condition. |
| Set Pulse Width | 20ns to 10.0sec | With Set Pulse Width highlighted, set the pulse width. |
| Polarity | Positive Negative | Select to trigger on positive or negative pulses. |
| Mode | Auto Normal | Select the type of trigger. The Normal mode is best for most pulse width trigger applications. |
| Coupling | AC DC HF Reject LF Reject | Select the components of the trigger signal applied to the trigger circuit. |
| More |  | Switch between submenu pages. |

Trigger When: The pulse width of the source must be ≥5ns so that the oscilloscope can detect the pulse.

＝, ≠: Within a ±5% tolerance, triggers the oscilloscope when the signal pulse width is equal to or not equal to the specified pulse width.

＜, ＞: Triggers the oscilloscope when the source signal pulse width is less than or greater than the specified pulse width.

Slope Trigger: Judges trigger according to the time for rising or falling, more flexible and accurate than the Edge trigger.

| Options | Settings | Comments |
| --- | --- | --- |
| Slope |  |  |
| Source | CH1 CH2 | Select the input source as the trigger signal. |
| Slope | Rising Falling | Select the slope type of signal. |
| Mode | Auto Normal | Select the type of trigger. The Normal mode is best for most pulse width trigger applications. |
| Coupling | AC DC Noise Reject HF Reject LF Reject | Selects the components of the trigger signal applied to the trigger circuitry. |
| Next Page |  |  |
| Vertical | V1 V2 | Adjust the vertical window by setting two trigger levels. Select this option and press F3 to choose V1 or V2. |
| When | = ≠ < > | Select the trigger condition. |
| Time | 20ns to 10.0sec | With this option highlighted, set the time span. |

Swap Trigger: As a feature of analog oscilloscopes, it gives stable displays of signals at two different frequencies. Mainly it uses a specific frequency to switch between two analog channels CH1 and CH2 so that the channels will generate swap trigger signals through the trigger circuitry.

| Options | Settings | Comments |
| --- | --- | --- |
| Swap Trigger |  |  |
| Mode | Auto Normal | Select the type of trigger. |
| Channel | CH1 CH2 | Push an option such as CH1, select the channel trigger type and set the menu interface. |
| Below list options in submenus. Swap Trigger allows CH1 and CH2 to select different trigger modes and to display waveforms on a same screen. That is, both channels can choose the following four trigger modes. |
| Type | Edge |  |
| Slope | Rising Falling |  |
| Coupling | AC DC HF Reject LF Reject | select the components of the trigger signal applied to the trigger circuitry. |
| Type | Video |  |
| Polarity | Normal Inverted |  |
| Standard | NTSC PAL/SECAM |  |
| Sync | All Lines Line Number Odd Field Even Field All Fields |  |
| Type | Pulse |  |
| Polarity | Positive Negative |  |
| When | = ≠ < > |  |
| Set Pulse Width | Pulse Width | Set the pulse width. |
| Coupling | AC DC Noise Reject HF Reject LF Reject |  |
| Type | Slope |  |
| Slope | Rising Falling | Select the slope type of signal. |
| Mode | Auto Normal | Select the type of trigger. The Normal mode is best for most pulse width trigger applications. |
| Coupling | AC DC Noise Reject HF Reject LF Reject | Selects the components of the trigger signal applied to the trigger circuitry. |
| Next Page |  |  |
| Vertical | V1 V2 | Adjust the vertical window by setting two trigger levels. Select this option and press F3 to choose V1 or V2. |
| When | = ≠ < > | Select the trigger condition. |
| Time | 20ns to 10.0sec | Set the time span. |

Overtime Trigger: In Pulse Width trigger, you may sometimes be puzzled with the long time for trigger, as you do not need a complete pulse width to trigger the oscilloscope, but want the trigger occurs just upon the overtime point. This is called Overtime Trigger.

| Options | Settings | Comments |
| --- | --- | --- |
| Type | OT |  |
| Source | CH1 CH2 | Select the trigger source. |
| Polarity | Positive Negative | Select to trigger on positive or negative pulses. |
| Mode | Auto Normal |  |
| Overtime | t |  |
| Coupling | AC DC HF Reject LF Reject | Selects the components of the trigger signal applied to the trigger circuitry. |

Holdoff: To use Trigger Holdoff, push the HORI button and set the Holdoff Time option. The Trigger Holdoff function can be used to generate a stable display of complex waveforms (such as pulse trains). Holdoff is the time between when the oscilloscope detects one trigger and when it is ready to detect another. During the holdoff time, the oscilloscope will not trigger. For a pulse train, the holdoff time can be adjusted to let the oscilloscope trigger only on the first pulse in the train.

## Menu and Option Buttons

As shown below, these four buttons at the front panel are used mainly to recall relative setup menus.

SAVE/RECALL: Displays the Save/Recall menu for setups and waveforms.

MEASURE: Displays the Measure menu.

ACQUIRE: Displays the Acquire menu.

UTIILITY: Displays the Utility menu.

CURSOR: Displays the Cursor menu.

DISPLAY: Displays the Display menu.

### SAVE/RECALL

Press the SAVE/RECALL button to save or recall oscilloscope setups or waveforms.

The first page shows the following menu.

| Options | Settings | Comments |
| --- | --- | --- |
| Waveforms |  |  |
| Source | CH1 CH2 off MATH off | Select a waveform display to store. |
| REF | RefA RefB | Select the reference location to store or recall a waveform. |
| Operation | Save | Save the source waveform to the selected reference location. |
|  | Ref on Ref off | Display or remove the reference waveform on the screen. |

Press ‘Next Page’ to recall the following menu.

| Options | Settings | Comments |
| --- | --- | --- |
| Setups |  |  |
| Operation Source | Flash memory USB disk | Store the current setups to the USB disk or the memory of the oscilloscope. |
| Memory | 0 to 9 | Specify the memory location in which to store the current waveform settings or from which to recall the waveform settings. |
| Operation | Save | Complete the saving operation. |
|  | Recall | Recall the oscilloscope settings stored in the location selected in the Setup field. Push the Default Setup button to initialize the oscilloscope to a known setup. |

See below for waveform menus.

Note: The oscilloscope will save the current settings if you wait 5 seconds after the last modification, and it will recall these settings the next time you power on the oscilloscope.

### MEASURE

Push the MEAS button to perform auto measurements. There are 11 types of measurements and up to 8 can be displayed at a time.

Click the MEAS key then the following menu appears.

| Options | Settings | Comments |
| --- | --- | --- |
| Source | CH1 CH2 | Select the measure source. |
| Measurement Type | Frequency | Calculate the waveform frequency by measuring the first cycle. |
|  | Period | Calculate the time of the first cycle. |
|  | Mean | Calculate the arithmetic mean voltage over the entire record. |
|  | Pk-Pk | Calculate the absolute difference between the greatest and the smallest peaks of the entire waveform. |
|  | Cyc RMS | Calculate the actual RMS measurement of the first complete cycle of the waveform. |
|  | Min | Examine the waveform record of all points in the current window and display the minimum value. |
|  | Max | Examine the waveform record of all points in the current window and display the maximum value. |
|  | Rise Time | Measure the time between 10% and 90% of the first rising edge of the waveform. |
|  | Fall Time | Measure the time between 90% and 10% of the first falling edge of the waveform. |
|  | Positive Width | Measure the time between the first rising edge and the next falling edge at the waveform 50% level. |
|  | Negative Width | Measure the time between the first falling edge and the next rising edge at the waveform 50% level. |
|  | Off | Do not take any measurement. |

Taking Measurements: For a single waveform (or a waveform divided among multiple waveforms), up to 8 automatic measurements can be displayed at a time. The waveform channel must stay in an ‘ON’ (displayed) state to facilitate the measurement. The automatic measurement can not be performed on reference or math waveforms, or in XY or Scan mode.

### ACQUIRE

Push the ACQUIRE button to set the acquisition parameter.

| Options | Settings | Comments |
| --- | --- | --- |
| Category | Real Time Equ-Time | Acquire waveforms by real-time digital technique. Rebuild waveforms by equivalent sample technique. |
| Mode (Real Time) | Normal Peak Detect Average | Acquire and accurately display most waveforms. Detect glitches and eliminate the possibility of aliasing. Reduce random or uncorrelated noise in signal display. The number of averages is selectable. |
| Averages (Real Time) | 4 16 64 128 | Select the number of averages. |
| Memory Depth (Real Time) | 4K, 40K, 512K, 1M | Select the memory depth for different board models. |

Normal: For the oscilloscope model with the bandwidth of 100MHz, the maximum sample rate is 1GS/s. For time base with insufficient sample rate, you may use the Sine Interpolation Algorithm to interpolate points between sampled points to produce a complete waveform record (4K by default).

Peak Detect: Use this mode to detect glitches within 10ns and to limit the possibility of aliasing. This mode is valid at the TIME/DIV setting of 4µs/div or slower. Once you set the TIME/DIV setting to 4µs/div or faster, the acquisition mode will change to Normal because the sample rate is fast enough that Peak Detect is unnecessary. The oscilloscope does not display a message to tell you that the mode has been changed to Normal.

Average: Use this mode to reduce random or uncorrelated noise in the signal to be displayed. Acquire data in Normal mode and then average a great number of waveforms. Choose the number of acquisitions (4, 16, 64 or 128) to average for the waveform.

Stopping the Acquisition: When you are running the acquisition, the waveform display is live. Stop the acquisition (press the RUN/STOP button) to freeze the display. In either mode, the waveform display can be scaled or positioned by vertical and horizontal controls.

Equivalent Acquisition: Just repeat the Normal acquisition. Use this mode to take a specific observation on repeatedly displayed periodic signals. You can get a resolution of 40ps, i.e. 25GSa/s sample rate, which is much higher than that obtained in real-time acquisition.

The acquisition principle is as follows.

As shown above, acquire input signals (cycle repeatable) for more than once at a slow sample rate, arrange the sample points by the time they appear, then recover waveforms.

### UTILITY

Push the UTILITY button to display the Utility Menu as follows.

| Options | Comments |
| --- | --- |
| System Info | Display the software and hardware versions, serial number and some other information about the oscilloscope. |
| Update Program | Insert a USB disk with upgrade program and the disk icon at the top left corner is highlighted. Press the Update Program button and the Software Upgrade dialog pops up. |
| Save Waveform | Insert a USB disk and the disk icon at the top left corner is highlighted. Click on this icon and you can see the waveform pause for a while, being saved. You can find the saved waveform data under the Hantek_x folder in the USB disk. Here X represents how many times you press the button. Each press generates a corresponding folder. For example, press once and a folder Hantek_1 is generated; press twice and two folders Hantek_1, Hantek_2 are generated. |
| Self Calibration | Press this option and the Self Calibration dialog pops up. Press F6 to perform the self calibration. |
| Advance | Buzzer and time setups See the dialog for buzzer and time setups |

Self Calibration: The self calibration routine can optimize the precision of the oscilloscope to fit the ambient temperature. To maximize the precision, you should perform the self calibration once the ambient temperature changes by 5℃ or more. Follow the instructions on the screen.

Tip: Press any menu button on the front panel to remove the status display and enter a corresponding menu.

### CURSOR

The Cursor Menu.

| Options | Settings | Comments |
| --- | --- | --- |
| Type | Off Voltage Time | Select a measurement cursor and display it. Voltage measures amplitude while Time measures time and frequency. |
| Source | CH1 CH2 MATH REFA REFB | Select a waveform to take the cursor measurement. Use the readouts to show the measurement. |
| Select Cursor | S E | S indicates Cursor 1. E indicates Cursor 2. A selected cursor is highlighted, which can be moved freely. Both cursors can be selected and moved at the same time. The box behind the cursor displays the location of the cursor. |
| Delta | Display the difference (delta) between the cursors. | Display the measurement in the box under this option. |

Moving Cursors: Press the key near Select Cursor to select a cursor and move it. Cursors can be moved only when the Cursor Menu is displayed.

### DISPLAY

The waveform display is affected by settings of the oscilloscope. A waveform can be measured once it is captured. The different styles to display a waveform on the screen give significant information about it.

There are two modes to display waveforms Single-window and Dual-window. Refer to Horizontal Controls for more information.

The DISPLAY menu.

| Options | Settings | Comments |
| --- | --- | --- |
| Type | Vectors Dots | Vectors fills up the space between adjacent sample points in the display; Dots only displays the sample points. |
| Persist | OFF 0.2S-8S selectable Infinite | Sets the time length to display each displayed sample point. |
| Format | YT XY | YT format shows the vertical voltage in relation to time (horizontal scale); XY format displays a dot between CH1 and CH2 each time when a sample is acquired, where the voltage or current of CH1 determines the X coordinate of the dot (horizontal) and the voltage or current of CH2 determines the Y coordinate (vertical). For detailed information, refer to descriptions on XY format in the following text. |
| Contrast |  | 0-15 16 ranks adjustable, with a progress bar to display. |
| Next Page |
| Grid | Dotted line Real line OFF | Off only displays the horizontal and vertical coordinates at the center graticule on the screen. |
| Grid Intensity |  | 0-15 16 ranks adjustable, with a progress bar to display. |

## Fast Action Buttons

AUTO: Automatically set the oscilloscope controls to generate a usable display of the input signals. Refer to the following table for relative content.

RUN/STOP: Continuously acquire waveforms or stop the acquisition.

### AUTO

Autoset is one of the advantages digital oscilloscopes have. When you push the AUTO button, the oscilloscope will identify the type of waveform (sine or square wave) and adjust controls according to input signals so that it can accurately display the waveform of the input signal.

| Functions | Settings |
| --- | --- |
| Acquire Mode | Adjusted to Normal or Peak Detect |
| Cursor | Off |
| Display Format | Set to YT |
| Display Type | Set to Vectors for an FFT spectrum; otherwise, unchanged |
| Horizontal Position | Adjusted |
| TIME/DIV | Adjusted |
| Trigger Coupling | Adjusted to DC, Noise Reject or HF Reject |
| Trigger Holdoff | Minimum |
| Trigger Level | Set to 50% |
| Trigger Mode | Auto |
| Trigger Source | Adjusted |
| Trigger Slope | Adjusted |
| Trigger Type | Edge |
| Trigger Video Sync | Adjusted |
| Trigger Video Standard | Adjusted |
| Vertical Bandwidth | Full |
| Vertical Coupling | DC (if GND was chosen before); AC for the video signal; otherwise, unchanged |
| VOLTS | Adjusted |

The Autoset function examines all channels for signals and displays corresponding waveforms. Autoset determines the trigger source according to the following conditions.

  - If multiply channels get signals, the oscilloscope will use the channel with the lowest frequency signal as the trigger source.

  - If no signals are found, the oscilloscope will use the lowest-numbered channel displayed in Autoset as the trigger source.

  - If no signals are found and no channels are displayed, the oscilloscope will display and use Channel 1 as the trigger source.

Sine Wave:

When you use the Autoset function and the oscilloscope determines that the signal is similar to a sine wave, the oscilloscope displays the following options.

| Sine Wave Options | Details |
| --- | --- |
| Multi-cycle Sine | Display multiple cycles that have appropriate vertical and horizontal scales. |
| Single-cycle Sine | Set the horizontal scale to display about one cycle of the waveform. |
| FFT | Convert the input time-domain signal to its frequency components and display the result as a graph of frequency versus amplitude (spectrum). Since it is a math calculation, see Section 5.3.1 Math FFT for more information. |
| Undo Setup | Let the oscilloscope recall the previous setup. |

Square Wave or Pulse:

When you use the Autoset function and the oscilloscope determines that the signal is similar to a square wave or pulse, the oscilloscope displays the following options.

| Square Wave Options | Details |
| --- | --- |
| Multi-cycle Square | Display multiple cycles that have appropriate vertical and horizontal scales. |
| Single-cycle Square | Set the horizontal scale to display about one cycle of the waveform. The oscilloscope displays Min, Mean and Positive Width automatic measurements. |
| Rising Edge | Display the rising edge. |
| Falling Edge | Display the falling edge. |
| Undo Setup | Let the oscilloscope recall the previous setup. |

## Signal Connectors

See the figure below to find the two signals connectors and a pair of metal electrodes at the bottom of the oscilloscope panel.

- CH1, CH2: Input connectors for waveform display, through which to connect and input the signal to be measured.

- Probe Compensation: Voltage probe compensation output and ground, used to electrically match the probe to the oscilloscope input circuit. The probe compensation ground and BNC shields connect to earth ground and are considered to be ground terminals. To avoid damages, do not connect a voltage source to any of these ground terminals.

# Application Examples

This chapter unfolds a further description on main features of the oscilloscope by giving eleven simplified application examples for reference to help solve your own test problems.

- Taking simple measurements

Using AUTO

Using the Measure menu to take auto measurements

- Taking cursor measurements

Measuring ring frequency and ring amplitude

Measuring pulse width

Measuring rise time

- Analyzing input signals to eliminate random noise

Observing a noisy signal

Eliminating random noise

- Capturing a single-shot signal

- Using X-Y mode

- Triggering on a pulse width

- Triggering on a video signal

Observing triggers on video fields and video lines

- Using Slope Trigger to capture a particular slope signal

- Using Overtime Trigger to measure a long pulse signal

- Using math functions to analyze waveforms

- Measuring data propagation delay

## Example 1: Taking Simple Measurements

When you want to observe an unknown signal in a certain circuit without having its amplitude and frequency parameters, you may use this function to take a fast measurement on the frequency, period and peak-to-peak amplitude of the signal.

Follow the steps below.

- Set the switch on the oscilloscope probe to 10X;

- Push the CH1 MENU button and set the Probe option attenuation to 10X;

- Connect the CH1 probe to the test point of the circuit;

- Press the AUTO button.

The oscilloscope will automatically set the waveform to a best display. If you want to further optimize the waveform display, you may manually adjust the vertical and horizontal controls until the waveform meets your particular requirement.

Taking Auto Measurements

The oscilloscope can display most signals by automatic measurements. To measure such parameters as the signal frequency, period, peak-to-peak amplitude, rise time and positive width, follow the steps below.

- Push the MEAS button to see the Measure menu.

- Select the first ‘unspecified’ option (marked by red arrow), and then enter the submenu.

- Select CH1 for the Source option. Then repeatedly select measure items in the Type menu. Push the back menu to go back to the measure interface. The corresponding box under the measure item shows the measurements.

- Repeat Step 2 and Step 3. Then select other measure items. Totally 8 measure items can be displayed.

Note: All readouts change with the measured signals.

The figure below shows three measure items as an example. The boxes under them display the measurements in large fonts.

## Example 2: Taking Cursor Measurements

You can use the cursor to quickly measure the time and amplitude of a waveform.

Measuring Ring Time (convertible to Frequency) and Amplitude on Rising Edge of Pulse

To measure the ring time on the rising edge of the pulse, follow the steps below.

- Push the CURSOR button to view the Cursor menu.

- Push F1 the Type option button and select Time.

- Push F2 the Source option button and select CH1.

- Select a cursor. If S is selected,move Cursor S on the screen; if E is selected, move Cursor E; if both are selected, move them at the same time.

- Put Cursor S on the first peak of the ring.

- Put Cursor E on the second peak of the ring.

- At Delta displays the measured time and at Cursor S an Cursor E display the positions of these two cursors.

- Push the Type option button and select Voltage.

- Put Cursor S on the highest peak of the ring.

- Put Cursor E on the lowest point of the ring. The amplitude of the ring will be displayed at Delta.

See figures below for better understanding.

Measuring Pulse Width

To analyze a pulse signal and to know its width, follow the steps below.

- Push the CURSOR button to view the Cursor menu.

- Push F1 the Type option button and select Time.

- Push F2 the Source option button and select CH1.

- Select a cursor. If S is selected, move Cursor S on the screen; if E is selected, move Cursor E; if both are selected, move them at the same time.

- Place Cursor S on the rising edge of the pulse and Cursor E on the falling edge.

- Thus at Delta displays the measured time and at Cursor S and Cursor E displays the time relative to the trigger.

See the figure below for better understanding.

Measuring Rise Time of Pulse

You may need to measure the rise time of the pulse in many application environments, usually, to measure the rise time between the 10% and 90% levels of the pulse waveform. To do so, follow the steps below.

- Click the TIME/DIV key to display the rising edge of the waveform.

- Click the VOLTS and Vertical Position keys to adjust the waveform amplitude to about 5 divisions.

- Push the CH1 MENU button.

- Push the VOLTS option button and select Fine. Click the Vertical Position key to accurately separate the waveform into 5 divisions.

- Click the Vertical Positon key to center the waveform. Position the waveform baseline to 2.5 divisions below the center graticule.

- Press the CURSOR button.

- Push the Type option button and select Time. Push the Source option button to select CH1.

- Select Cursor S and place it at the 10% level of the waveform.

- Select Cursor E and place it at the 90% level of the waveform.

- The Delta readout in the Cursor Menu is the rise time of the pulse.

See the figure below for better understanding.

## Example 3: Analyzing Input Signals to Eliminate Random Noise

In certain circumstances, to display a noisy signal on the oscilloscope and to get its details, you may follow the steps below to analyze this signal.

Observing Noisy Signal

- Enter the Acquire menu.

- Push the Type option button and select Real Time.

- Push the Peak Detect option button.

- If necessary, enter the DISPLAY menu and set the Contrast option to view the noise more clearly.

See the figure below for better understanding.

Eliminating Random Noise

- Enter the Acquire menu.

- Push the Type option button and select Real Time.

- Push the Average option button.

- Push the Averages option button and adjust the number of running averages to watch the change in the waveform display.

Note: Averaging reduces random noise and let you view the signal details more easily.

See the figure below for better understanding

## Example 4: Capturing Single-shot Signal

You may refer to the following example to easily capture some aperiodic signals like pulses and glitches.

To set for a single-shot acquisition, follow the steps below.

- First, set up the oscilloscope probe and the attenuation factor of CH1.

- Click the vertical VOLTS and horizontal TIME/DIV keys to a proper position for a better examination of the signal.

- Enter the Acquire menu.

- Push the Peak Detect option button.

- Push the TRIG MENU button and select Rising for the Slope option. Then adjust the trigger level properly.

- Select Single mode to start the acquisition.

Using this feature can help you to capture occasional events more easily. This is an advantage of the HandHeld oscilloscope.

## Example 5: Using X-Y Mode

Viewing Phase Differences between Two Channel Signals

For example, you need to measure the change in a phase across a circuit network.

Connect the oscilloscope with circuitry and view the input and output of the circuit in XY mode. Follow the steps below.

- First, prepare two oscilloscope probes and set the switches to 10X on both probes.

- Push the CH1 MENU button and set the Probe option attenuation to 10X; push the CH2 MENU button and set the Probe option attenuation to 10X.

- Connect the CH1 probe to the input of the network, and connect the CH2 probe to the output.

- Push the AUTO button.

- Click the VOLTS keys to display approximately the same amplitude signals on each channel.

- Enter the Display menu.

- Push the Format option button and select XY.

- Now the oscilloscope displays a Lissajous pattern to characterize the input and output of the circuit.

- Click the VOLTS and Vertical Position keys to properly scale the waveform display.

- Use the Lissajous's oscillographic method to observe and calculate the phase differences following the formula below.

As sinθ=A/B or C/D, in which θ is the phase difference angle between channels and A, B, C, D represent what shown in the figure below, you can get the value of the phase difference angle by the formula: θ=±arcsin(A/B) or ±arcsin(C/D).

If the principal axes of the ellipse are in the first and third quadrants, the phase difference angle should be in the first and fourth quadrants, i.e. within (0~π/2) or (3π/2~2π). If the principal axes of the ellipse are in the second and fourth quadrants, the phase difference angle should be in the second and third quadrants, i.e. within (π/2~π) or (π-3π/2). See the figure below for better understanding.

## Example 6: Triggering on Pulse Width

Triggering on a Specific Pulse Width

While testing the pulse width of a signal in a circuit, you may need to verify the pulse width is consistent with the theoretic value. Or even if the edge triggering shows that your signal has the same pulse width with the specific signal, you still doubt about the result. Then you can follow the steps below.

- Set the Probe option attenuation to 10X.

- Push the AUTO button to trigger a stable waveform display.

- Push the Single Cycle option button in the Autoset menu and read out the signal pulse width.

- Push the TRIG MENU button.

- Select Pulse for the Type option; Select CH1 for the Source option; Click the TRIGGER LEVEL key to set the trigger level at the bottom of the signal.

- Select the When option button and select ‘=’.

- Push the Set Pulse Width option button. Set the pulse width to the value read out in Step 3.

- Click the TRIGGER LEVEL key to set the pulse width to the value read out in Step 3.

- Push the More option button and select Normal for the Mode option. Once triggering on normal pulses, the oscilloscope can give a stable waveform display.

- When option is set to >, < or ≠ and there appear any aberrant pulses that meet the specified condition, the oscilloscope will trigger. For example, the signal contains such aberrant pulses as shown below, you may select ‘≠’ or ‘<’ to trigger on the pulse.

As shown in the above figure, you can get a stable waveform display if inputting a square wave at the frequency of 1KHz, with pulse width set to 500μs.

## Example 7: Triggering on Video Signal

Assume that you are monitoring the video signals of a television to see if they are input normally, and the video signal is of an NTSC system. You can get a stable display by using the video trigger.

Triggering on Video Fields

To trigger on the video fields, follow the steps below.

- Push the TRIG MENU button to see the Trigger menu.

- Push F1 to select Video for the Type option.

- Push the Source option button to select CH1; push the Polarity option button to select Normal; push the Standard option button to select NTSC.

- Push the Sync option button to select Odd Field, Even Field or All Fields.

- Push the Trigger Level button to adjust the trigger level and stabilize video signals.

- Push the horizontal TIME/DIV and the Vertical Position keys to display on the screen a complete video signal triggering on a video field.

The figure below shows a stable signal triggering on a video field.

Triggering on Video Lines

To trigger on the video lines, follow the steps below.

- Push the TRIG MENU button to see the Trigger menu.

- Select Video for the Type option.

- Push the Source option button to select CH1; push the Polarity option button to select Normal; push the Standard option button to select NTSC; push the Sync option button to select Line Number.

- Click the Trigger Level key to adjust the trigger level and stabilize video signals.

- Adjust the line number (NTSC: 0-525 lines).

- Click the horizontal TIME/DIV and the vertical VOLTS keys to display on the screen a complete video signal triggering on a video line. See the figure below.

## Example 8: Using Slope Trigger to Capture Particular Slope Signal

In many occasions, we are not only concerned about the edge of the signal, but also want to know the rise and fall times of the signal. To better observe this kind of signals, we bring in the slope trigger. Follow the steps below.

- Push the TRIG MENU button to see the Trigger menu.

- Select Slope for the Type option.

- Push the Source option button to select CH1; push the Slope option button to select Rising; push the Mode option button to select Auto; push the Coupling option button to select DC.

- Click the ‘Next Page’ button and select Vertical. Adjust V1 and V2 to proper locations. Select the When option button and set it to ‘=’.

- Select ‘Time’ and adjust the time until you get a stable display of waveforms. See the figure below.

## Example 9: Using Overtime Trigger to Measure Long Pulse Signal

It is not easy to observe some part of a long pulse signal by using the edge or pulse width trigger. In such case, you can use the overtime trigger by following steps.

- Push the TRIG MENU button to see the Trigger menu.

- Select OT for the Type option; push the Polarity option button to select Normal; push the Mode option button to select Auto; push the Coupling option button to select DC.

- Click the Trigger Level key to adjust the trigger level and stabilize video signals.

- Adjust the line number (NTSC: 0-525 lines).

- Clck the horizontal TIME/DIV and the vertical VOLTS keys to display on the screen a complete video signal triggering on a video line. See the figure below.

Note: The difference between the overtime and the delay triggers is that the overtime trigger can identify the pulse you need according to your set time and trigger on any point of the pulse. In the other word, the overtime trigger occurs based on pulse identification. It is similar to the > mode of the pulse width trigger, but not the same.

## Example 10: Using Math Functions to Analyze Waveforms

Using math functions to analyze input waveforms is another advantage of the digital oscilloscope. For example, you want to get the instantaneous difference between two channel waveforms. By using the math function of the oscilloscope, you can get a better representation of the waveform on the screen. To observe this signal, follow the steps below.

- Set the Probe option attenuation to 10X.

- Open CH1 and CH2 at the same time, both with the attenuation of 10X.

- Push the AUTO button to trigger a stable waveform.

- Push the M/R MENU button to see the Math menu.

- Push the Operation option button and select ‘CH1+CH2’.

- Click the horizontal TIME/DIV and the vertical VOLTS keys to properly scale the waveform for easy check.

In addition, the oscilloscope also supports the - and FFT functions. For a detailed analysis on FFT, refer to Chapter 5.3.1 Math FFT.

Note: You should compensate both probes before performing the math operation; otherwise, differences in probe compensation will appear as errors in the differential signal.

As illustrated in the above figure, input a 1KHz sine wave from CH1 and a 1KHz square wave from CH2.

Follow the above steps to set up the Math menu, and observe the subtracted waveform as shown in the figure below.

Those in pink are added waveforms.

## Example 11: Measuring Data Propagation Delay

When you doubt that there appear instabilities in a serial data propagation control circuit, you can set the oscilloscope to measure the propagation delay between the enable signal and the transfer data.

To set the propagation delay measurement, follow the steps below.

- Connect two oscilloscope probes respectively to the CS (chip-select) pin and the DATA pin on the chip.

- Set the Probe option attenuation to 10X for both probes.

- Open CH1 and CH2 at the same time, both with the attenuation of 10X.

- Push the AUTO button to trigger a stable waveform display.

- Adjust the horizontal and vertical controls to optimize the waveform display.

- Push the CURSOR button to view the Cursor menu.

- Push the Type option button and select Time.

- Select Cursor S and place it on the active edge of the enable signal.

- Select Cursor E and place it on the data output transition (See the figure below).

- Read the data propagation delay in the Delta readout.

## Example 12: Setting Net IP

To set net IP, follow the steps below.

1. Select the menu Utility->System->Pc Set->Mode->NET.Then there pushes up the NetCfg dialog.

2. Select Device IP or Device Port you want to set and press ENTER key.Then there pushes up the Input panel dialog.

3. Modify the parameter by the soft keyboard.

# Multimeter

About this chapter

This chapter provides an introduction to the multimeter functions of DSO1000B Series. The introduction gives guides to show how to use the menus and perform basic measurements.

Connecting the Meter

Use the 4-mm safety banana jack inputs for the Meter functions: 10A , mA, COM,V/Ω/C.

Multimeter Operation Window

Figure 7-1 Multimeter operation window

Description

1) Measurement mode indictors:

DC: Direct electric measurement

AC: Alternating electric measurement

2)The symbol of Mutimeter current mode.

3) Manual/Auto range indictors, among which the MANUAL refers to measuring range in manual

operation mode and Auto means the measuring range in automatic operation mode.

4) The reading value of measurement.

5) The bar graph indictor.

6) DC or AC Measurement mode control.

7) Absolute /relative magnitude measuring control: The sign “||” expresses the absolute magnitude

measuring control and “ ” represents the relative magnitude measuring control.

8) Manually or automatically measuring range control.

Operating the Multimeter

If you are in the oscilloscope window, press OSC/DMM key, the oscilloscope will switch to the

multimeter mode window.Then the screen will display the measure mode window that was in use

the last time before you quit multimeter measure. When you switch to the multimeter measurement

the first time, the default measure mode is DC voltage mode.

Measuring Resistance Values

To measure a resistance, do the following:

1) Press the R key and then resistance measurement window appears on the screen.

2) Insert the black lead into the COM banana jack input and the red lead into the V/O/C banana jack

input.

3) Connect the red and black test leads to the resistor. The resistance value is shown on the screen

in Ohm.

Then, the screen will look like the following figure 7-2.

Figure 7-2 Resistance Measurement

Making a Diode Measurement

To make a measurement on the diode, do the following:

1) Press the diode key and a diode symbol appears at the top of the screen. .

2) Insert the black lead into the COM banana jack input and the red lead into the V/O/C banana jack

input.

3) Connect the red and black leads to the diode and the voltage value of the diode is displayed on

the screen in volt.

Then, the screen will look like the following figure 7-3.

Figure 7-3 Diode Measurement

Making an On-off Measurement

To perform an On-off test, do the following:

1) Press the On-off key and then On-off indictor appears on the top of the screen.

2) Insert the black lead into the COM banana jack input and the red lead into the V/O/C banana jack

input.

3) Connect the red and black leads to the tested points. If the resistance value of the tested points

is less than 30 O, you will hear beep sound from the test tool.

Then, the screen will look like the following figure 7-4.

Figure 7-4 On-off Measurement

Making a Capacitance Measurement

To measure a capacitance, do the following:

1) Press the C key and a capacitor symbol appears on the top of the screen.

2) Insert the black lead into the COM banana jack input and the red lead into the V/O/C banana jack

input.

3) Connect the red and black leads to the capacitor and the capacitance value is displayed on the

screen in µF or nF.

Then, the screen will look like the following figure 7-5.

Figure 7-5 Capacitance Measurement

Making a DC Voltage Measurement

To measure a DC voltage, do the following:

1. Press the V key and DC appears at the top of the screen.

2. Insert the black lead into the COM banana jack input and the red lead into the V/O/C banana jack

input.

3.Connect the red and black leads to the measured points and the voltage value of measured

points is displayed on the screen.

Then ,the screen will look like the following figure 7-6.

Figure 7-6 DC voltage Measurement

Making an AC Voltage Measurement

To measure the AC voltage, do the following:

1) Press the V key and DC appears on the screen.

2) Press the F1 key and AC appears on the screen.

3) Insert the black lead into the COM banana jack input and the red lead into the V/O/C banana jack

input.

4) Connect the red and black leads to the measured points and the AC voltage value of measured

points will be displayed on the screen.

Then, the screen will look like the following figure 7-7.

Figure 7-7 AC voltage Measurement

Making a DC Current Measurement

To measure a DC current which is less than 600 mA , do the following:

1) Press the A key and then DC appears on the screen. The unit on the main reading screen is mA.

Press F2 to switch the measurement between mA and 10A. 600mA is acquiescently.

2) Insert the black lead into the COM banana jack input and the red lead into the mA banana jack

input.

3) Connect the red and black leads to the measured points and the DC current value of measured

points will be displayed on the screen.

Then, the screen will look like the following figure 7-8.

Figure 7-8 DC current Measurement for 600 Measurement

To measure a DC current which is larger than 600 mA, do the following:

1) Press the A key and then DC appears on the screen. The unit on the main reading screen is mA.

2) Press F2 key to switch to 10A measurement, the unit on the main reading screen is A.

3) Insert the black lead into the COM banana jack input and the red lead into the 10A banana jack

input.

4) Connect the red and black leads to the measured points and the DC current value of the

measured points will be displayed on the screen.

5) Press F2 to return to 600 mA measurement.

Then, the screen will look like the following figure 7-9.

Figure 7-9 DC Current Measurement for 10A

Making an AC Current Measurement

To measure an AC current which is less than 600 mA, do the following:

1) Press the A key and then DC appears on the screen. The unit on the main reading screen is

mA,and mA will display on the bottom of the screen, press F2 to switch the measurement between

mA and 10A. 600mA is acquiescently.

2) Press the F1 key once and AC will display on the bottom of the screen.

3) Insert the black lead into the COM banana jack input and the red lead into the mA banana jack

input.

4) Connect the red and black leads to the measured points and the AC current value of measured

points will be displayed on the screen.

Then, the screen will look like the following figure 7-10.

Figure 7-10 AC Current Measurement for 600 mA

To measure an AC current which is larger than 600 mA, do the following:

1) Press the A key and then DC appears on the screen. The unit on the main reading screen is mA.

2) Press F2 key to switch to10A measurement, the unit on the main reading screen is A.

3) Press the F1 key once and AC will display on the bottom of the screen.

4) Insert the black lead into the COM banana jack input and the red lead into the 10A banana jack

input.

5) Connect the red and black leads to the measured points and the AC current value of the

measured points will be displayed on the screen.

6) Press F2 to return to 600 mA measurement.

Then, the screen will look like the following figure 7-11.

Figure 7-11 AC Current Measurement for 10A

Taking a Relative Measurement

A currently measured result relative to the defined reference value is displayed in a relative

measurement.

The following example shows how to take a relative measurement. At first, it is required to

acquire a reference value.

1) Press the Okey.

2) Insert the black lead into the COM banana jack input and the red lead into the V/O/C banana jack

input.

3) Connect the red and black test leads to the resistor. The resistance value is shown on the screen

in Ohm.

4) When the reading leveling off, press F1 key and then ||/△is displayed on the top of the screen.

The saved reference value is displayed beside.

Then, the screen will look like the following figure 7-12.

Figure 7-12 Relative Measurement

Selecting Automatic/Manual Range Adjustment

The default range mode of the instrument is automatic range. Suppose you are using the

DC voltage mode, to switch to the manual range, perform the following steps:

1) Press F3 key to enter the manual range mode and then Manual is displayed on the top of the

screen.

2) Under the manual range mode, the measuring range is increased by a stage when pressing F4

key each time, and when reaching the highest stage, it jumps to the

lowest stage by pressing F4 key once again.

3) Press F3 key to switch back to the automatic range mode and then Auto is displayed on the top

of the screen.

Attention: capacitance measurement without manual range mode.

Then,the screen will look like the following figure 7-13.

Figure 7-13 The Manual Range Mode

# Troubleshooting

## Problem Settlement

- If the oscilloscope does not start up at power on, follow these steps:

- Check the power cord to verify it has been connected properly;

- Check the power on/off button to ensure it has been pushed;

- Then restart the oscilloscope.

Contact your local HANTEK distributor or directly keep touch with HANTEK Technical Support department if the oscilloscope still can not be turned on normally.

- If there is no display of waveforms on the screen when the oscilloscope is turned on, follow these steps:

- Check the probe to assure its proper connection to the input BNC;

- Check the channel switch (such as CH1, CH2 menu buttons) to make sure it has been turned on;

- Check the input signal to verify it has been connected to the probe correctly;

- Affirm that all measured circuits have signals to output;

- Turn up the magnitude for DC signals with large magnitude;

- In addition, you may press the Auto Measure button to perform an automatic detection of signals at first.

Contact HANTEK Technical Support department in time if there is still no display of waveforms.

- If the waveform of the input signal is distorted seriously, follow these steps:

- Check the probe to assure its proper connection to the channel BNC;

- Check the probe to assure its good connection to the measured object;

- Check the probe to verify it has been well calibrated. Otherwise, refer to the content about calibration described in this manual.

- If the waveform is rolling continuously on the screen but can not be triggered, follow these steps:

- Check the trigger source to make sure it consistent with the input channel;

- Check the trigger level to assure its correct adjustment.

- Check the trigger mode to confirm it is a right choice for the input signal. The default trigger mode is edge trigger. However, it is not suitable for all kinds of input signals.

# Specifications

## Technical Specifications

All specifications herein mentioned apply to the DSO1000B series oscilloscopes. Before checking an oscilloscope from HANTEK to see if it complies with these specifications, make sure it meets the following conditions:

- The oscilloscope must have been operating continuously for twenty minutes under the specified operating temperature.

- The Do Self Cal operation must be performed through the Utility menu if the operating temperature changes by more than 5℃.

- The oscilloscope must be within the factory calibration interval.

All specifications are guaranteed unless noted ‘typical’.

Oscilloscope Specifications

Horizontal

| Sample Rate Range | 1GS/s |
| --- | --- |
| Waveform Interpolation | (sin x)/x |
| Record Length | Maximum 1M samples per single-channel; maximum 512K samples per dual-channel (4K,16K,40K optional) |
| TIME/DIV Range | DSO1062B DSO1102B | DSO1202B |
|  | 4ns/div to 40s/div, in a 2, 4, 8 sequence | 2ns/div to 40s/div, in a 2, 4, 8 sequence |
| Sample Rate and Delay Time Accuracy | ±50ppm over any ≥1ms time interval |
| Delta Time Measurement Accuracy (Full Bandwidth) | Single-shot, Normal mode ± (1 sample interval +100ppm × reading + 0.6ns) |
|  | >16 averages ± (1 sample interval + 100ppm × reading + 0.4ns) |
|  | Sample interval = s/div ÷ 200 |
| Position Range | DSO1062B DSO1102B |
|  | 4ns/div to 8ns/div | (-8div × s/div) to 20ms |
|  | 20ns/div to 80μs/div | (-8div × s/div) to 40ms |
|  | 200μs/div to 40s/div | (-8div × s/div) to 400s |
|  | DSO1202B |
|  | 2ns/div to 10ns/div | (-4div × s/div) to 20ms |

Vertical

| A/D Converter | 8-bit resolution, each channel sampled simultaneously |
| --- | --- |
| VOLTS Range | 2mV/div to 5V/div at input BNC |
| Position Range | 2mV/div to 200mV/div, ±2V >200mV/div to 5V/div, ±50V |
| Analog Bandwidth in Normal and Average modes at BNC or with probe, DC Coupled | 2mV/div to 20mV/div, ±400mV 50mV/div to 200mV/div, ±2V 500mV/div to 2V/div, ±40V 5V/div, ±50V |
| Selectable Analog Bandwidth Limit, typical | 20MHz |
| Low Frequency Response (-3db) | ≤10Hz at BNC |
| Rise Time at BNC, typical | DSO1062B | DSO1102B | DSO1202B |
|  | <5.8ns | ＜3.5ns | <1.8ns |
| DC Gain Accuracy | ±3% for Normal or Average acquisition mode, 5V/div to 10mV/div ±4% for Normal or Average acquisition mode, 5mV/div to 2mV/div |
| DC Measurement Accuracy, Average Acquisition Mode | Measurement Type: Average of ≥16 waveforms with vertical position at zero Accuracy: ± (3% × reading + 0.1div + 1mV) when 10mV/div or greater is selected |
|  | Measurement Type: Average of ≥16 waveforms with vertical position not at zero Accuracy: ± [3% × (reading + vertical position) + 1% of vertical position + 0.2div] Add 2mV for settings from 2mV/div to 200mV/div; add 50mV for settings from 200mV/div to 5V/div |
| Volts Measurement Repeatability, Average Acquisition Mode | Delta volts between any two averages of ≥16 waveforms acquired under same setup and ambient conditions |

Note: Bandwidth reduced to 6MHz when using a 1X probe.

Trigger

| Trigger Sensitivity (Edge Trigger Type) | Coupling | Sensitivity |
| --- | --- | --- |
|  | DC | Source | DSO1062B DSO1102B | DSO1202B |
|  |  | CH1 CH2 | 1div from DC to 10MHz; 1.5div from 10MHz to Full | 1.5div from 10MHz to 100MHz; 2div from 100MHz to Full |
|  | AC | Attenuates signals below 10Hz |
|  | HF Reject | Attenuates signals above 80kHz |
|  | LF Reject | Same as the DC-coupled limits for frequencies above 150kHz; attenuates signals below 150kHz |
| Trigger Level Range | Source | Range |
|  | CH1, CH2 | ±8 divisions from center of screen |
| Trigger Level Accuracy, typical (Accuracy is for signals having rise and fall times ≥20ns) | Source | Accuracy |
|  | CH1、CH2 | 0.2div × volts/div within ±4 divisions from center of screen |
| Set Level to 50%, typical | Operates with input signals ≥50Hz |

Note: Bandwidth reduced to 6MHz when using a 1X probe.

| Video Trigger Type | Source | Range |
| --- | --- | --- |
|  | CH1, CH2 | Peak-to-peak amplitude of 2 divisions |
| Signal Formats and Field Rates, Video Trigger Type | Supports NTSC, PAL and SECAM broadcast systems for any field or any line |  |
| Holdoff Range | 100ns to 10s |  |

| Pulse Width Trigger |  |
| --- | --- |
| Pulse Width Trigger Mode | Trigger when < (Less than), > (Greater than), = (Equal), or ≠ (Not Equal); Positive pulse or Negative pulse |
| Pulse Width Trigger Point | Equal: The oscilloscope triggers when the trailing edge of the pulse crosses the trigger level. Not Equal: If the pulse is narrower than the specified width, the trigger point is the trailing edge. Otherwise, the oscilloscope triggers when a pulse continues longer than the time specified as the Pulse Width. Less than: The trigger point is the trailing edge. Greater than (also called overtime trigger): The oscilloscope triggers when a pulse continues longer than the time specified as the Pulse Width. |
| Pulse Width Range | Selectable from 20ns to 10s |

| Slope Trigger |  |
| --- | --- |
| Slope Trigger Mode | Trigger when < (Less than), > (Greater than), = (Equal), or ≠ (Not Equal); Positive slope or Negative slope |
| Slope Trigger Point | Equal: The oscilloscope triggers when the waveform slope is equal to the set slope. Not Equal: The oscilloscope triggers when the waveform slope is not equal to the set slope. Less than: The oscilloscope triggers when the waveform slope is less than the set slope. Greater than: The oscilloscope triggers when the waveform slope is greater than the set slope. |
| Time Range | Selectable from 20ns to 10s |
| Overtime Trigger | The leading edge: Rising edge or Falling edge; Time Setting: 20-10s |

| Swap Trigger |  |
| --- | --- |
| CH1 | Internal Trigger: Edge, Pulse Width, Video, Slope |
| CH2 | Internal Trigger: Edge, Pulse Width, Video, Slope |

| Trigger Frequency Counter |  |
| --- | --- |
| Readout Resolution | 6 digits |
| Accuracy (typical) | ±30ppm (including all frequency reference errors and ±1 count errors) |
| Frequency Range | AC coupled, from 4Hz minimum to rated bandwidth |
| Signal Source | Pulse Width or Edge Trigger modes: all available trigger sources The Frequency Counter measures trigger source at all times, including when the oscilloscope acquisition pauses due to changes in the run status, or acquisition of a single shot event has completed. Pulse Width Trigger mode: The oscilloscope counts pulses of significant magnitude inside the 1s measurement window that qualify as triggerable events, such as narrow pulses in a PWM pulse train if set to < mode and the width is set to a relatively small time. Edge Trigger mode: The oscilloscope counts all edges of sufficient magnitude and correct polarity. Video Trigger mode: The Frequency Counter does not work. |

Acquisition

| Acquisition Modes | Normal, Peak Detect, and Average |
| --- | --- |
| Acquisition Rate, typical | Up to 2000 waveforms per second per channel (Normal acquisition mode, no measurement) |
| Single Sequence | Acquisition Mode | Acquisition Stop Time |
|  | Normal, Peak Detect | Upon single acquisition on all channels simultaneously |
|  | Average | After N acquisitions on all channels simultaneously, N can be set to 4, 8, 16, 32, 64 or 128 |

Inputs

| Inputs |  |
| --- | --- |
| Input Coupling | DC, AC or GND |
| Input Impedance, DC coupled | 1MΩ±2% in parallel with 20pF±3pF |
| Probe Attenuation | 1X, 10X |
| Supported Probe Attenuation Factors | 1X, 10X, 100X, 1000X |
| Maximum Input Voltage | Overvoltage Category | Maximum Voltage |
|  | CAT I and CAT II | 300VRMS (10×), Installation Category |
|  | CAT III | 150VRMS (1×) |
|  | Installation Category II: derate at 20dB/decade above 100kHz to 13V peak AC at 3MHz* and above. For non-sinusoidal waveforms, peak value must be less than 450V. Excursion above 300V should be of less than 100ms duration. RMS signal level including all DC components removed through AC coupling must be limited to 300V. If these values are exceeded, damage to the oscilloscope may occur. |

Measurements

| Cursors | Voltage difference between cursors: △V Time difference between cursors: △T Reciprocal of △T in Hertz (1/ΔT) |
| --- | --- |
| Automatic Measurements | Frequency, Period, Mean, Peak-to-peak, Cycle RMS, Minimum, Maximum, Rise Time, Fall Time, Positive Width, Negative Width |

General Specifications

| Display |
| --- |
| Display Type | 5.7 Inch width TFT Display |
| Display Resolution | 480 (Vertical) X 640(Horizontal) pixels |
| Display Contrast | Adjustable (16 gears) with the progress bar |
| Probe Compensator Output |
| Output Voltage, typical | About 5Vpp into ≥1MΩ load |
| Frequency, typical | 1kHz |
| Power Supply |
| Switching Adatper | AC Input:100-240VACRMS,0.6A MAX,50Hz-60Hz; DC Output:9V,2A |
| DC Input | DC8.5-15V,2A |
| Power Consumption | <30W |
| Environmental |
| Temperature | Operating: 32℉ to 122℉ (0℃ to 50℃) |
|  | Nonoperating: -40℉ to 159.8℉ (-40℃ to +71℃) |
| Cooling Method | Convection |
| Humidity | +104℉ or below (+40℃ or below): ≤90% relative humidity |
|  | 106℉ to 122℉ (+41℃ to 50℃): ≤60% relative humidity |
| Altitude | Operating and Nonoperating | 3,000m (10,000 feet) |
|  | Random Vibration | 0.31gRMS from 50Hz to 500Hz, 10 minutes on each axis |
|  | Nonoperating | 2.46gRMS from 5Hz to 500Hz, 10 minutes on each axis |
| Mechanical Shock | Operating | 50g, 11ms, half sine |
| Mechanical |
| Size | Length | 245mm |
|  | Height | 163mm |
|  | Depth | 52mm |
| Weight |  | 1.2 Kg |

Meter Mode

| Maximum Resolution | 6000 Counts |
| --- | --- |
| DMM Testing Modes | Voltage,Current,Resistance,Capacitance,Diode & Continuity |
| Maximum Input Voltage | AC : 600V DC : 800V |
| Maximum Input Current | AC : 10A DC : 10A |
| Input Impedance | 10MΩ |

Meter Specification

| Range | Accuracy | Resolution |
| --- | --- | --- |
| DC Voltage | 60.00mV(manual) | ±1%±1digit | 10uV |
|  | 600.0mV |  | 100uV |
|  | 6.000V |  | 1mV |
|  | 60.00V |  | 10mV |
|  | 600.0V |  | 100mV |
|  | 800V |  | 1V |
| AC Voltage | 60.00mV(manual) | ±1%±3digit | 10uV |
|  | 600.0mV(manual) |  | 100uV |
|  | 6.000V |  | 1mV |
|  | 60.00V |  | 10mV |
|  | 600.0V |  | 100mV |
| DC Current | 60.00mA | ±1.5%±1digit | 10uA |
|  | 600.0mA | ±1%±1digit | 100uA |
|  | 6.000A | ±1.5%±3digit | 1mA |
|  | 10.00A |  | 10mA |
| AC Current | 60.00mA | ±1.5%±3digit | 10uA |
|  | 600.0mA | ±1%±1digit | 100uA |
|  | 6.000A | ±1.5%±3digit | 1mA |
|  | 10.00A |  | 10mA |
| Resistance | 600.0 | ±1%±1digit | 0.1Ω |
|  | 6.000K |  | 1Ω |
|  | 60.00K |  | 10Ω |
|  | 600.0K |  | 100Ω |
|  | 6.000M |  | 1KΩ |
|  | 60.00M | ±1.5%±3digit | 10KΩ |
| Capacitance | 40.00nF | ±1%±1digit | 10pF |
|  | 400.0nF |  | 100pF |
|  | 4.000uF |  | 1nF |
|  | 40.00uF |  | 10nF |
|  | 400.0uF |  | 100nF |
|  | Attention:The smallest capacitance value that can be measured is 5nF. |
| Diode | 0V~2.0V |
| On-off Test | < 30Ω |

## Accessories

All the following accessories are available by contacting your local HANTEK distributor.

Standard Accessories

| Sketch | Description |
| --- | --- |
|  | X1, X10 two passive probes. The passive probes have a 6MHz bandwidth (rated 100Vrms CAT III) when the switch is in the X1 position, and a maximum bandwidth (rated 300Vrms CAT II) when the switch is in the X10 position. Each probe consists of all necessary fittings. |
|  | A quick guide of the DSO1000B series oscilloscopes. It gives a description mainly on functions and operating basis of these oscilloscopes. |
|  | A Power Adapter that fits the standard of destination country. |

ptional Accessories

| Sketch | Description |
| --- | --- |
|  | A USB A-B line, used to connect external devices with USB-B interface like a printer or to establish communications between PC and the oscilloscope. |
|  | A software installation CD. It contains the user manual of DSO1000B, giving particular descriptions on the DSO1000B series oscilloscopes. |

# Services and Support

Thank you for choosing HANTEK. Please contact us through the following ways should you have any inquiry regarding our products. We shall do our best to help you.

- Contact your local HANTEK distributor;

- Contact your local HANTEK field office;

- Contact HANTEK headquarters in China.

Headquarters

Qingdao Hantek Electronic Co., Ltd

http://www.hantek.com

Address: 5/F, No.177 Zhuzhou Road (Huite Industry City), Qingdao, China

266101

Tel: +86-532-88703687 / 88703697

Fax: +86-532-88705691

Email: service@hantek.com.cn

Technical Support

Tel: +86-532-88703687 (ext: 607)

Email: support@hantek.com.cn

Marketing Department

Tel: +86-532-88703687 (ext: 603)

Email: david@hantek.com.cn

Sales Department

Tel: +86-532-88703687 (ext: 604)

Email: rose@hantek.com.cn

# General Care and Cleaning

## General Care

Do not put or leave the device in a place where the LCD display will be exposed to direct sunlight for long periods of time.

Note: To avoid damage to the oscilloscope or probes, do not expose them to sprays, liquids, or solvents.

## Cleaning

Examine the oscilloscope and probes as often as operating conditions require. To clean the exterior surface, perform the following steps:

- Use a lint-free cloth to remove floating dust on the outside of the oscilloscope and probes. Take care to avoid scratching the glabrous display filter.

- Use a soft cloth dampened with water to clean the oscilloscope. For more efficient cleaning, you may use an aqueous solution of 75% isopropyl alcohol.

Note: To avoid damage to the surface of the oscilloscope or probes, do not use any corrosive or chemical cleaning agents.

# Appendix A  Harmful and Poisonous Substances or Elements

|  | Harmful and poisonous substances or elements 1 |
| --- | --- |
| Component2 | Pb | Hg | Cd | Cr(Vi) | PBB | PBDE |
| Shell and Chassis | X | 0 | 0 | X | 0 | 0 |
| Display Module | X | X | 0 | 0 | 0 | 0 |
| Circuit Board | X | 0 | 0 | X | 0 | 0 |
| Power Supply | X | 0 | 0 | X | 0 | 0 |
| Electric Wire and Cable Assembly | X | 0 | 0 | 0 | 0 | 0 |
| Connector | X | 0 | 0 | X | 0 | 0 |
| Fastener and Installed Hardware | X | 0 | X | X | 0 | 0 |
| Other Accessories (including probes) | X | 0 | 0 | X | 0 | 0 |
| Others | 0 | 0 | 0 | 0 | 0 | 0 |

‘X’ means that at least the content of this poisonous and harmful substance in a homogeneous material of this component exceeds the limit specified in the SJ/T 11363-2006 standard.

‘0’ indicates that the content of this poisonous and harmful substance in all homogeneous materials of this component is refrained under the limit stated in the SJ/T 11363-2006 standard.

This component list contains components approved in the file ‘Management Measures’.

# Text boxes

Cursor

Trigger slope can be rising or falling

Measurement

Input Terminal

CAUTION

Refer to Manual

DATA

Measurement

Ground Terminal

CS

CH1

17719

1062B/1102B/1202B

DATA

Major Window

Trigger level can be adjusted vertically

Cursor

CS

CH2

Signal Horizontal Centering

D

C

B

A

5 divisions

Compensated correctly

Overcompensated

Undercompensated

Voltage Cursor

Time Cursor

Fourth Acquisition

Third Acquisition

Mains Connected

ON (Power)

Second Acquisition

First Acquisition

DSO1000B SERIES

HANDHELD OSCILLOSCOPE

9

1    2 3    4           5           6   7       8

Input repeated signals

Normal Mode Acquires a Single Sample Point in Each Interval

10

9

1

6

: XY

: Gray indicates auto persistence; Green means persistence display is enabled. When the icon is set to green, the time for persistence display will be shown behind it.

: Vectors

: YT

5

Protective Ground (Earth) Terminal

4

2

7

3

8

Normal Acquisition Intervals

Sample Points

The readouts in big font size on the menu are just results of the corresponding measurements.

At most 9 groups of setups can be stored

The white waveforms on the menu is the recalled RefA waveform

Acquisition Interval

Acquisition Interval

Holdoff

Holdoff

Indicates Trigger Points

Trigger Level

Triggers when pulse is less than width setting

: Dots

Location of expanded window data in memory

Minor Window (Expanded Window)

Threshold level

Tolerance

Tolerance

Triggers when pulse is not equal to width setting ±5%

Falling Edge

Rising Edge

Triggers when pulse is greater than width setting

Threshold level

High Voltage

168

= Trigger Point

Triggers when pulse is equal to width setting ±5%

14   15

13 12  11

10

Amplitude Cursors

Frequency Cursors

PROBE COMP

1

CH1: to connect with the probe

5

4

3

2

Frequency component

Mains Disconnected

OFF (Power)

USER’S MANUAL

Fundamental frequency component