USER’S MANUAL

1. SUMMARIZE

This meter is a handheld automatic range digital multiemter which driven by battery. It can be used to measure such parameters as DCV,  ACV, ACA and DCA, Resistance, Continuity Buzzer, Diode, Frequency etc.

2. GENERAL FEATURE

Displaying：Max. displaying:4000，

Sampling rate: approx. 3 times/second.

Polarity indication：Automatic negative polarity indication

Over range indication：” OL” shown on the LCD

Low battery indication:“      ” appears

Operation environment：0℃～40℃，R.H.< 75%.

Storage environment：-10℃～50℃，R.H.<80%.

Power supply：（Size AAA battery）1.5V×2

Size：193×73×37mm

Weight：about 220g.

Accessories：user’s manual，battery, probes.

3. TECHNICAL CHARACTERISTIC

Accuracy calibration  Temperature:(23±5)℃，R.H.<75%

3-1.ACV（V～）

| RANGE | ACCURACY | RESOLUTION |
| --- | --- | --- |
| 4V | ±(1.0%+10d) | 1mV |
| 40V |  | 10mV |
| 400V |  | 100mV |
| 600V | ±(1.0%+6d) | 1V |

Input resistance：10MΩ

Frequency response：Sin wave and triangular wave is 40Hz～1kHz

3-2.DCV（    ）

| RANGE | ACCURACY | RESOLUTION |
| --- | --- | --- |
| 400mV | ±(0.5%+7d) | 0.1mV |
| 4V |  | 1mV |
| 40V |  | 10mV |
| 400V |  | 100mV |
| 600V | ±(1.0%+6d) | 1V |

Input resistance：10MΩ

3-3.RESISTANCE（auto range）

| RANGE | ACCURACY | RESOLUTION |
| --- | --- | --- |
| 400Ω | ±(1.0%+5d) | 0.1Ω |
| 4KΩ | ±(1.0%+5d) | 1Ω |
| 40KΩ |  | 10Ω |
| 400KΩ |  | 100Ω |
| 4MΩ |  | 10kΩ |
| 40MΩ | ±(1.2%+10d) | 100kΩ |

3-4.ACA  A～(auto range）

| RANGE | ACCURACY | RESOLUTION |
| --- | --- | --- |
| 40A | ±(4.0%+10d) | 0.001A |
| 400A |  | 0.01A |
| 600A |  | 0.1A |

Frequency response：50～60Hz

3-5.DCA  A～(auto range）

| RANGE | ACCURACY | RESOLUTION |
| --- | --- | --- |
| 40A | ±(4.0%+10d) | 0.001A |
| 400A |  | 0.01A |
| 600A |  | 0.1A |

3-6.FREQUENCY  Hz(auto range）

| RANGE | ACCURACY |
| --- | --- |
| 0～10MHz | ±(0.3%+3d) |

Shows 0 when the signal is less than 3Hz；sensitivity：1V RMS

3-7. CAPACITANCE (auto range)

| RANGE | ACCURACY | RESOLUTION |
| --- | --- | --- |
| 4nF | ±(5.0%+40d) | 1pF |
| 40nF | ±(5.0%+10d) | 10pF |
| 400nF |  | 100pF |
| 4uF |  | 1nF |
| 40uF |  | 10nF |
| 1000uF |  | 1uF |

4. KEY PANEL DESCRIPTION

4-1.Data Hold Key（HOLD/BL）

When press the key, the “HOLD” symbol will be shown on the LCD and the last reading of the meter, until press the key again or turn the function switch, the symbol disappear and the readings comes back. Press “HOLD/BL”key to turn on the meter, the “APO” symbol disappear, the auto power off function canceled. If turn on the meter by switch like usual, the auto power off function is on, the meter will turn off in 15 minutes if without any operation. This key got a backlight function, press this key for 3 seconds, the backlight on; if press this key again for more than 3 seconds, the backlight off.

4-2.Select Key（SELECT）

Turn switch to     or     range, press the key can choose the function you wish.

4-3. ACA And DCA Reset Key (REL)

Turn switch to the      range, press the select key to get the measurement function of ACA or DCA. If the LCD not shows 00.00, press the “REL” key to clear. The LED screen will display the REL symbol, reset again after short-press, exit after long-press for 3 seconds.

5.OPERATE INSTRUCTIONS

5-1.ACV/DCV measurement

Turn switch to ACV or DCV range

a. Connect the red test lead to VΩ terminal, and black test lead to COM terminal.

b. Connect the test leads to the circuit which should be tested. Get the result directly. (When it is under the DCV, the polarity of the red test lead will be shown together on the LCD.)

5-2.Resistance Measurement

a. Turn to range Ω.

b. Connect the red test lead to VΩ terminal, and black test lead to COM terminal.

c. Connect the test leads to the circuit or resistance which should be tested.

NOTE:

- To measure resistance which is more than 2MΩ, the meter needs few seconds to get the stable readings, and for the high resistance, this is a normal situation.

- When it is out of connection or under open circuit, it will shows “OL”.

- When the online resistor is measured, be sure to shut off the line power supply and discharge all capacitors completely.

5-3.Diode and Continuity test（buzzer function）

a. Connect the red test lead to VΩ terminal, and black test lead to COM terminal.

b. Turn to range      , and press “SELECT” to choose the range diode or continuity buzzer.

c. If the resistance between both ends of the circuit is less than about 30Ω, the built-in buzzer will sound.

d. Under the range DIODE, connect the red test probe onto the positive of the diode and the black test probe onto the negative. The forward conduction voltage drop will be shown on the LCD.

5-4.ACA/DCA measurement

Turn switch to the      range

press the select key to get the measurement function of ACA or DCA. If the LCD not shows 00.00, press the “REL” key to clear. The LED screen will display the REL symbol, reset again after short-press, exit after long-press for 3 seconds.

Place the wire vertically in the center of the clamp head, and the reading is the ACA or DCA value.

Note: Do not press REL key during measurement

5-6.Frequency Measurement（Hz）

a. Connect the red test lead to VΩ terminal, and black test lead to COM terminal.

b. Turn to range Hz, Connect the test leads to the circuit which should be tested.

5-7.Capacitance Measurement

a. Connect the red test lead to VΩ terminal, and black test lead to COM terminal.

b. Turn to range      , (note that the polarity of the red test lead is positive.)

c. Connect the test leads to the circuit which should be tested. And check if the polarity is correct.

NOTE：be sure to discharge all capacitors completely before measurement, when testing the big capacitance, it needs about 10 seconds to get the stable value.

5-8.NCVmeasurement

Turn to range NCV, connect the clamp head to the wire or socket, when the wire or socket got any AC signal through, the meter will judge the signal and buzzer in different sound.

6.BATTERY REPLACEMENT

When        symbol shows on the LCD, please replace the battery (AAA  1.5V).

7.Note

In order to avoid the damage of the meter, please do not input 250V DC or RMS AC under range       /Ω/Hz/      .